## Protocols
Folder Name | Protocol Name | Range Name | Recon Name | Hidden Protocol | Hidden Range | CT Scan&GO Approved | Protocol Description | RT Patient Marking | Patient Type | Organ Characteristic | Target Region | FAST Topo Region | Scan Mode | Accessible Scan Modes | Strategy Review | Decision Trees | Scan Strategies | Tilt [°] | Acquisition [mm] | Scan Direction | Table Direction | Patient Position | Feed [mm] | No. of Scans | Tube Position | Length [mm] | CARE Dose4D & CARE keV | CARE Dose4D | CARE keV IQ level | CARE keV optimized for | CARE keV Grouping | kV | kV (L/H) | Quality ref. mAs | Quality ref. mAs (L) | Quality ref. mAs (H) | Eff. mAs | Eff. mAs (L) | Eff. mAs (H) | mA | CARE keV min. kV | CARE keV max. kV | CTDIvol [mGy] | SSDE [mGy] | CTDIvol [mGy/s] | DLP [mGy*cm] | DLP [mGy*cm/s] | Dose Notification CTDIvol [mGy] | Dose Notification CTDIvol [mGy/s] | Dose Notification DLP [mGy*cm] | Dose Notification DLP [mGy*cm/s] | Energy Thresholds [keV] | Shaped Filter | HandCARE | mA average | Pitch | Rot. Time [s] | Scan Time [s] | Exam Time [s] | Cycle Time [s] | Patient Instruction | Instruction Command | Post Instruction Delay [s] | Absolute Start Delay: Actual ¦ Intended [s] | Post Bolus Delay (BT): Actual ¦ Intended [s] | Post Bolus Delay (TB) [s] | FAST ROI | Trigger Level @120kV [HU] | Auto Trigger | Exposure Pattern [s] | Memo | Check&GO Metal Detection | Check&GO | Check&GO Image Type | Check&GO Orientation | Scan Start | Scan Ref. Lines | FAST Adjust Upper Limit Exam Time [s] | FAST Adjust Lower Limit max. mAs [%] | FAST Adjust Lower Limit Pitch Factor | Pulsing Mode | Phase Start | Phase End | Phase Unit | Est. Heart Rate | Est. Respiratory Rate | Temporal Resolution [ms] | Injector Coupling | Contrast Protocol | Slice Thickness [mm] | Increment [mm] | Slices/rot. (IVR) | No. of Images | Comment 1 | Comment 2 | Reconstruction Pattern [s] | Recon Strategies | Spectral Recon | Auto keV | Energy [keV] | Inline Results | Multi Recons | Fix Axial | Image Type | Window@120kV | FAST Window | Kernel | QIR | QIR Strength | iBHC | iBHC Material | iMAR | iMAR Type | Image Orientation | Image Order | FAST Planning | FAST Planning Width | FAST 3D | Matrix Size | Matrix | Width [mm] | Height [mm] | Center X [mm] | Center Y [mm] | Mirroring | HD FoV | Recon Type | Recon Phase Step | Recon Phase(s) | Recon Phase Unit | StackRecon | Temporal Resolution [ms] | Flex ECG Recon | ECG Documentation | Auto Recon | Recon Ref. Lines | Body Part Examined | Series Splitting | Transfer Nodes | Auto Transfer | Workplace | Auto CT View&GO | CT View&GO Dataroles | Auto Printing | Auto Printing Layout | Advanced Visualization | Advanced Visualization Dataroles | Local Application Workflow | syngo.via Workflow | syngo.via Datarole | Inline Results | Filtering of Lung CAD Markers | Start Position [mm] | End Position [mm] | CPR Type | Vessel | Vessel Label | Organ | Angle Between Images [°] | Start Angle [°] | End Angle [°] | VRT Preset | Remove Table | Remove Bones | Viewing Direction | High Resolution | Application Class | Auto keV | Energy [keV] | Show Plaques | Data Set | Scoring Mode | Reference Database | Structure Template Selection Mode | Structure Templates | Lung Analysis View | Thick Axial | Lung Lobe Segmentation | Unfolding Type
Customer Protocols
Head | 1.01 Hjärna (N) (Adult) |  |  | no |  | yes | Vuxenprotokoll.
|  | Topogram | [6 empty] | Adult |  | Head | Head | ScanplanningTopoAdultQuantumSn | ScanplanningTopoAdultQuantumSn | Off |  |  |  | 36 x 0,20 | Head to Feet | In | Head First - Supine (HFS) |  |  | Left | 256 | [5 empty] | Sn 140 | [7 empty] | 55 |  |  | SPR (16 cm) 0,13 |  |  | 3,19 |  | 0,00 |  | 0,0 |  |  | Standard |  | 55,0 |  |  |  | 1,57 |  | None |  |  | 2 | [7 empty] | On |  |  |  | START key
|  |  | Topogram 0,20 Tx20 sag | [101 empty] | Topogram Tx20 Neuro |  | Tx20 | [29 empty] | Head |  | PACS | On | Acquisition Workplace | Off |  | Off
|  | Topogram | [6 empty] | Adult |  | Head | Head | ScanplanningTopoAdultQuantumSn | ScanplanningTopoAdultQuantumSn | Off |  |  |  | 36 x 0,20 | Feet to Head | Out | Head First - Supine (HFS) |  |  | Bottom | 256 | [5 empty] | Sn 140 | [7 empty] | 55 |  |  | SPR (16 cm) 0,13 |  |  | 3,19 |  | 0,00 |  | 0,0 |  |  | Standard |  | 55,0 |  |  |  | 1,57 |  | None |  |  | 2 | [7 empty] | On |  |  |  | START key
|  |  | Topogram 0,20 Tx20 cor | [101 empty] | Topogram Tx20 Neuro |  | Tx20 | [29 empty] | Head |  | PACS | On | Acquisition Workplace | Off |  | Off
|  | Hjärna_N | [6 empty] | Adult | Brain | Head |  | RoutineSpiralAdultHeadQuantumplus | RoutineSpiralAdultHeadQuantumplus; RoutineSpiralAdultHeadQuantumplusXcare; RoutineSequenceAdultHeadQuantumplus | Off |  |  |  | 144 x 0,40 | Feet to Head | Out | Head First - Supine (HFS) |  |  |  |  | Off |  |  |  |  | 120 |  |  |  |  | 228 | [5 empty] | (16 cm) 38,5 |  |  | 661 |  | 0,00 |  | 0,0 |  |  | Standard |  | 159,6 | 0,35 | 0,50 |  | 5,45 |  | None |  |  | 2 | [8 empty] | None |  |  | START key | Off | 8 | 25 | 0,10 | [7 empty] | Off
|  |  | Hjärna_N 0,80 Hr40 Q4 RECON FRÅN DENNA | [86 empty] | 0,80 | 0,40 |  | 405 |  |  |  |  | Monoenergetic Plus | Off | 67 | None | None | On |  | Hjärna W70 C35 | Off | Hr40 | On | 4 | On | Quantum | Off |  | Axial | Feet to Head | None |  |  | 512 | Square | 210,00 | 210,00 | 0,00 | 0,00 | None | Off | [8 empty] | On | Off | Head |  | PACS; SYNGOVIA_EC | On | Acquisition Workplace | Off | None | Off |  | None | None | None | MM Reading | AA - Hjärna_N Ax 4mm MPR AutoRR; AA - Hjärna_N Cor 4mm MPR AutoRR; AA - Hjärna_N Sag 4mm MPR AutoRR | None
|  |  | Hjärna_N 0,40 Hr72 Q2 | [86 empty] | 0,40 | 0,40 |  | 406 |  |  |  |  | Monoenergetic Plus | Off | 67 | None | None | On |  | Ben W3000 C1000 | Off | Hr72 | On | 2 | On | Quantum | Off |  | Axial | Feet to Head | None |  |  | 512 | Square | 210,00 | 210,00 | 0,00 | 0,00 | None | Off | [8 empty] | Off | Off | Head |  | PACS | On | Acquisition Workplace | Off | None | Off |  | None | None | None | MM Reading | Auto | None
|  |  | Hjärna_N 1,00 Qr44 Q4 VNC | [86 empty] | 1,00 | 0,70 |  | 231 |  |  |  |  | VNC |  |  | None | None | On |  | Neuro VNC W70 C35 | Off | Qr44 | On | 4 | On | Quantum | Off |  | Axial | Feet to Head | None |  |  | 512 | Square | 210,00 | 210,00 | 0,00 | 0,00 | None | Off | [8 empty] | Off | Off | Head |  | PACS | On | Acquisition Workplace | Off | None | Off |  | None | None | None | MM Reading | Auto | None
|  |  | Hjärna_N 1,00 Qr44 Q4 Iodine | [86 empty] | 1,00 | 0,70 |  | 231 |  |  |  |  | Iodine |  |  | None | None | On |  | Neuro Iodine W90 C30 | Off | Qr44 | On | 4 | On | Quantum | Off |  | Axial | Feet to Head | None |  |  | 512 | Square | 210,00 | 210,00 | 0,00 | 0,00 | None | Off | [8 empty] | Off | Off | Head |  | PACS | On | Acquisition Workplace | Off | None | Off |  | None | None | None | MM Reading | Auto | None